## Round 1
Testing Information |  | Test # | Resin | Dimensions | Puntured? | T1 | T2 | T3 | Avg Penetration Depth | 10% avg | Std Dev | Uniform Penetration? | Notes
Flow rate: 5.1 mL / hr |  | 1 | Rigid | 5 x 5 | yes | 0.3 | 0.3 | 0.3 | 0.3 | 0.03 | 0 | 1 |  | need zoomed in photo on
Testing time for each sample: 2 minutes (starting time when fluid reaches microneedle array) |  | 2 | Rigid | 5 x 5 | yes | 0.2 | 0.2 | 0.2 | 0.2 | 0.02 | 3.3993e-17 | 1 |  | missing photo on SD card
MNA held into porcine skin during the test |  | 3 | Clear | 3 x 3 | yes, completely penetrated skin sample | 1.3 | 0.6 | 0.1 | 0.66667 | 0.066667 | 0.60277 | 0 | Resistance when removing array from skin
Penetration depth measured at 3 locations using calipers |  | 4 | Clear | 4 x 4 | no |  |  |  |  | 0 | 0 | 0 | needle tips too blunt, some needle tips broken, loose inlet for tubing
After testing, excess ink on surface of skin blotted with paper towel |  | 5 | Rigid | 5 x 5 | yes | 0.1 | 0.2 | 0.1 | 0.13333 | 0.013333 | 0.057735 | 0 | slight resistance when removing microneedle array from skin
Uniform penetration defined as equal penetration depth at 3 measurement sites |  | 6 | Clear | 5 x 5 | yes | 0.3 | 0.2 | 0.3 | 0.26667 | 0.026667 | 0.057735 | 0
|  | 7 | Clear | 3 x 3 | yes | 0.9 | 0.3 | 0.1 | 0.43333 | 0.043333 | 0.41633 | 0 | needles were bending
|  | 8 | Surgical guide | 5 x 5 | yes | 1 | 0.9 | 0.9 | 0.93333 | 0.093333 | 0.057735 | 1
|  | 9 | Surgical guide | 10 x 10 | yes | 0.4 | 0.4 | 0.6 | 0.46667 | 0.046667 | 0.11547 | 0
| Trial | MNA Dimensions | Resin | Punctured skin? | Uniform Penetration?
| 1 | 5 x 5 | Rigid | Yes | Yes
| 2 | 5 x 5 | Rigid | Yes | Yes
| 3 | 3 x 3 | Clear | Yes | No
| 4 | 4 x 4 | Clear | No | No
| 5 | 5 x 5 | Rigid | Yes | No
| 6 | 5 x 5 | Clear | Yes | No
| 7 | 3 x 3 | Clear | Yes | No
| 8 | 5 x 5 | Surgical guide | Yes | Yes
| 9 | 10 x 10 | Surgical guide | Yes | No

## Round 2
MNA Penetration Testing Data
Testing Information | MNA # | Resin | Dimensions | Punctured? | Trial 1 | Trial 2 | Trial 3 | Avg Penetration Depth | 20% avg | Std Dev | Uniform Penetration?
Flow rate: 5.1 mL / hr | 1 | Rigid | 5 x 5 | Yes | 0.6 | 0.5 | 0.5 | 0.53333 | 0.10667 | 0.057735 | Yes
Testing time for each sample: 2 minutes (starting time when fluid reaches microneedle array) | 2 | Rigid | 5 x 5 | Yes | 0.4 | 0.7 | 0.5 | 0.53333 | 0.10667 | 0.15275 | No
MNA held into porcine skin during the test | 3 | Clear | 3 x 3 | Yes (completely punctured skin) | 0.8 | 0 | 0 | 0.26667 | 0.053333 | 0.46188 | No
Penetration depth measured at 3 locations using calipers | 4 | Clear | 4 x 4 | Yes | 0.4 | 0.2 | 1.2 | 0.6 | 0.12 | 0.52915 | No
After testing, excess ink on surface of skin blotted with paper towel | 5 | Rigid | 5 x 5 | Yes | 0.9 | 0.6 | 0.8 | 0.76667 | 0.15333 | 0.15275 | Yes
Uniform penetration defined as equal penetration depth at 3 measurement sites | 6 | Tough | 5 x 5 | Yes | 0.1 | 0.1 | 0.5 | 0.23333 | 0.046667 | 0.23094 | No
| 7 | Clear | 3 x 3 | Yes (completely punctured skin) | 0 | 0 | 0 | 0 | 0 | 0 | No
| 8 | Surgical Guide | 5 x 5 | Yes (completely punctured skin) | 1.6 | 1.3 | 1.4 | 1.4333 | 0.28667 | 0.15275 | Yes
| 9 | Surgical Guide | 10 x 10 | Yes | 1.4 | 0.6 | 1.2 | 1.0667 | 0.21333 | 0.41633 | No